UNITED STATES DEPARTMENT OF COMMERCE

National Oceanic and Atmospheric Administration

NATIONAL ENVIRONMENTAL SATELLITE, DATA

AND INFORMATION SERVICE

Silver Spring, MD  20910

September 15, 2015

To:

Arron Layns, Algorithm Management Project Manager

Lihang Zhou, JPSS STAR Algorithm Manager

Dr. Fuzhong Weng, OMPS SDR Team Lead

Dr. James Gleason, JPSS Project Scientist

From: Dr. Mitch Goldberg, JPSS Program Scientist

Re: Approval of OMPS Nadir Profiler and OMPS Total Column Earth View (EV) SDR as Validated

During the validate maturity reviews in December 2013, and the delta review in June 2014,  the algorithm maturity review board recommended two issues to be resolved before the OMPS (NP and TC) SDRs can be declared validated.   The two issues were to 1) Improve cross-track effects in NM (i.e. scan direction dependent error) and 2) Improve stray-light correction in NP.   For the first issue, the team:

- Updated wavelength LUT and solar flux LUT

- Validated wavelength dependent cross-track irradiance error from solar observation data

- Validated wavelength dependent cross-track normalized radiance error from SDR EV data via the TOMS RADiative transfer model (TOMRAD)

- Validated cross-track SO2 index variation from NM EDR data

The results now show much better cross-track agreement and meet the SDR requirement.     For the second issue, the team reevaluated the NP stray light correction and found the current stray light look-up table provides adequate calibration for the NP.  The team has also made other significant accomplishments including:

- Routine instrument and SDR performance monitoring, characterization, and improvement.

- Improvement of EV normalized radiance consistency between NP and NM by 2-10% in 300-310 nm.

- Instrument cal/val support (e.g., orbit adjustment, anomaly resolution)

- Complete documentation (Users’ Guide)

- Weekly dark LUTs update and delivery

- Completed Configuration Change Request (CCR) 1879 and 2100: NM  stray light LUT update (in operation Dec. 18, 2014) – Corrected out of range stray light

- Completed CCR 2546, (Discrepancy Report (DR) 7825 and 7826) modified CDFCB and xml files

- Completed CCR 2548 NP radiometric LUTs updated and delivered

- SDR software improved: stray light correction

In summary the OMPS Earth View (EV), SDRs meet the validated maturity:

- OMPS EV SDRs meet SDR performance requirement as well as EDR products requirement

  - The cross-track direction radiance accuracy meets spec and the  error is less than  2.0% with updated wavelength and day one solar LUTs

  - The NM and NP consistency in 300-310 nm has been improved by 2-10% with updated radiance calibration coefficients

  - Sensor orbital performance is stable and meet expectation

Both OMPS NM and NP EV SDR should be declared as validated-maturity products, effective August 20th 2015.  Congratulations to the OMPS SDR team - excellent work!!

Sincerely,

Dr. Mitchell D. Goldberg

JPSS Program Scientist

Algorithm Maturity Review Board Chair

Appendix: