Ozone Mapper Profile Suite (OMPS) Sensor Data Record (SDR)

Error Budget

OMPS SDR Science Team

September 08, 2015

# Geolocation

## Specification

The boresight alignment uncertainty between the nadir instrument interface and the nadir alignment reference shall be less than 160 arcsec (1/4 of a pexel) in both cross-track and along-track direction (1 sigma). The total cumulative boresight alignment shift between the final ground calibration and or-orbit operations shall be less than 500 arcsec (3 sigma, per axis) [1].

## Pre-Launch Budget

To be provided.

## Pre-Launch Performance

Alignment uncertainty was 90 arcsec per axis (to be verified). Alignment shift, estimated by analysis, was 95 arcsec in azimuth and 222 aresec in elevation.

## Post-Launch Performance

Less than 64 arcsec or 1/10 of a pixel.

Figure 1: High-Spatial-Resolution Geolocation Comparison. The image on the left shows a false color map of the OMPS effective reflectivity (from a single Ultraviolet channel at 380 nm) over the Arabian Peninsula region for January 30, 2012 when the instrument was making a special set of high-spatial-resolution measurements with 5×10 km2 FOVs at nadir. The color scale intervals range from 0 to 2 % in dark blue to 18 to 20 % in yellow. The image on the right is an Aqua Moderate Resolution Imaging Spectroradiometer (MODIS) Red-Green-Blue image for the same day.

# Signal-Noise-Ratio (SNR)

## Specification

The TC sensor shall meet or exceed the minimum signal-to-noise ratio (SNR) requirement OF 1000 for each wavelength channel over the range of input radiance levels (SS3.2.1.7.1-1) [1]. The NP sensor shall meet or exceed the minimum SNR requirement for each wavelength over the range of input radiance levels shown (SS3.2.1.7.1-3) [1].

## Pre-Launch Budget

Budget for this error is 1000 for NM and 45-400 channel dependent for NP.

## Pre-Launch Performance [2]

| Instrument | Mode | worst case SNR | Average SNR |
| --- | --- | --- | --- |
| NP | radiance | 2892 | 6055 |
|  | irradiance | 2204 | 4723 |
| TC | radiance | 3301 | 3547 |
|  | irradiance | 2698 | 2895 |

## Post-Launch Performance

Meet the specification with margins.

Figure 2: OMPS Nadir Mapper earth view noise as percent of the earth view signal. The root mean square residuals in percent were computed from an Empirical Orthogonal Function Analysis where the six largest patterns were removed.  For the NM sensor, shown on the left, the residuals equate to signal-to-noise ratios of better than 2000:1 for wavelengths greater than 320 nm, dropping to 1000:1 at the shorter wavelengths.  For the NP sensor, shown on the right, the SNR varies from 550:1 at the longer wavelength to 250:1 for the shortest; spikes correspond to lower signal levels due to solar lines at those wavelengths.

Figure 3: SNR from solar views (working diffuser) for OMPS NP (left) and NM (right).  The dotted red line in each figure is the required SNR.

# Albedo Calibration

## Specification

The wavelength independent radiometric (albedo) calibration uncertainty shall be less than 2%. The wavelength dependent radiometric (albedo) calibration uncertainty shall be less than 0.5% [2].

## Pre-Launch Error Budget

| Source of Uncertainty | Albedo 1σ Fractional Uncertainty (%) |  |  |  |
| --- | --- | --- | --- | --- |
|  | λ - independent |  | λ - dependent |  |
|  | NP | TC | NP | TC |
| Requirement | 2.0 | 2.0 | 0.5 | 0.5 |

## Pre-Launch Performance

| Source of Uncertainty | Albedo 1σ Fractional Uncertainty (%) |  |  |  |
| --- | --- | --- | --- | --- |
|  | λ - independent |  | λ - dependent |  |
|  | NP | TC | NP | TC |
| OMPS NPP RSS Total | 1.653 | 1.717 | 0.426 | 0.497 |

## Post-Launch Performance

| Source of Uncertainty | Albedo 1σ Fractional Uncertainty (%) |  |  |  |
| --- | --- | --- | --- | --- |
|  | λ - independent |  | λ - dependent |  |
|  | NP | TC | NP | TC |
| OMPS Percent Uncertainty | ≤ 2.0 | ≤ 2.0* | ≤ 0.5 | ≤ 0.5 |

# Figure 3 NM albedo uncertainty meets the requirement for the most of the channel. Exception is cross track position #35 on wavelength > 360 nm where the uncertainty varies from 2 - 3% as show in the following figure. Soft calibration on the EDR level is being implemented to eliminate this residual error.

# Nonlinearity

## Specification

The system nonlinearity shall be less than 2% full well. Uncertainty of system nonlinearity shall be less than 0.5% [2].

## Pre-Launch Budget

Budget for nonlinearity is 2%. Budget for nonlinearity uncertainty is 0.5%.

## Pre-Launch Performance

System nonlinearity is less than 2%. Uncertainty in nonlinearity is 0.1% [2].

## Post-Launch Performance

To be verified via EDR performance, including soft calibration.

Figure 4: Maximum nonlinearity on-orbit is ~0.4%, meeting the specification of 2%. Uncertainty in nonlinearity is 0.2%, meeting the specification of 0.5%.

# Sensor Response Uniformity (intra-IFOV)

## Specification

The accuracy of the calibration at one pixel with respect to another pixel should be 1% or better (SS3.2.1.19.1-1).

## Pre-Launch Budget

The accuracy of the calibration at one pixel with respect to another pixel should be 1% or better.

## Pre-Launch Performance

The sensor response uniformity (intra-IFOV) will not be directly verified by test. OMPS meet this requirement by using an integrating sphere, which produces an essentially flat field. That means that when two different pixels are calibrated, they look at the same source with the same intensity. The accuracy of the calibration with this source is driven by its stability. The integrating sphere stability is 0.6% peak-to-peak or better. Thus, the relative response knowledge is better than 1%.

## Post-Launch Performance

The sensor response uniformity (intra-IFOV) will not be directly verified by orbital data. A statistical method is being used to study the pixel-to-pixel variation. This method requires a large set of high resolution earth view data over the Greenland. This study is currently undergoing and more high resolution data is being collected.

# Stray Light

## Specification

Response to signal shall be less than < 2%.

## Pre-Launch Budget

Budget for this error is 2%.

## Pre-Launch Performance

A summary of stray light contributions to signal and overall uncertainty in radiance calibrations [2].

|  | NP |  | NM |  |
| --- | --- | --- | --- | --- |
|  | Stray Light (%) | Residual Uncertainty (%) | Stray Light (%) | Residual Uncertainty (%) |
| Radiance | 2.75 | 0.242 | 0.5 | 0.043 |

## Post-Launch Performance

Figure 5 Post-launch validation uses change in EV Mg II index that reflects stray light correction. Mg II index varied with SZA before the NP stray light correction. Stray light contamination caused up to -1.0 ~ 5.0% errors in EV radiance. The errors are within ±2% for the most of the channels after stray light correction was applied [3].

# Wavelength Registration

## Specification

Uncertainty of band center wavelength shall be less than 0.02 nm.

## Pre-Launch Budget

Budget for this error is 0.02 nm.

## Pre-Launch Performance

Ground wavelength calibration uncertainty is <0.01 nm (NDR3.2.1.2.3.1-4 [4]).

## Post-Launch Performance

# Figure 6 Orbital wavelength change is < ±0.01 nm for NM and ±0.02 nm for NP.

# Intra-Orbital Wavelength Stability

## Specification

Intra-orbital wavelength stability was not specified, but the EDR sensitivity expects this error to be less than 0.02 nm.

## Pre-Launch Budget

Budget for this error is 0.014 nm.

## Pre-Launch Performance

During the OMPS Integrated Sensor Suite (ISS) thermal balance testing, temperature cycles were imposed that are representative of variations expected on orbit, and the spectral shifts were recorded. 5% of these data were then used to train Singular Value Decomposition (SVD) model, which was then validated using the remaining 95% of data. The SVD model then predicts that on orbit shifts to be less than 0.01 nm.

## Post-Launch Performance

Figure 7 Intra-orbital wavelength shift is smaller than 0.04 pixels that is equivalent to 0.016 nm.

# Irradiance uncertainty

## Specification

The absolute accuracy for laboratory radiometric measurements between 250 nm and 380 nm shall be ≤7% for irradiance (SS3.2.1.11.1.1-1) [2].

| Source of Uncertainty | Irradiance Absolute 1σ Fractional Uncertainty (%) |  |
| --- | --- | --- |
|  | NP | NM |
| Requirement | 7.0 | 7.0 |

## Pre-Launch Budget

Budget for this error is 7% for both NP and NM [2].

## Pre-Launch Performance [2]

| Source of Uncertainty | Irradiance Absolute 1σ Fractional Uncertainty (%) |  |
| --- | --- | --- |
|  | NP | NM |
| OMPS NPP RSS Total | 3.499 | 3.194 |

## Post-Launch Performance [3]

| Source of Uncertainty | Irradiance Absolute 1σ Fractional Uncertainty (%) |  |
| --- | --- | --- |
|  | NP | NM |
| OMPS NPP RSS Total | 7.0* | 7.0 |

*see figure 8 for explanation.

Figure 8 Most channels meet the requirement. Exception is the NP noise channels 250-250 nm and 306-310 nm where uncertainty slightly exceeds the requirement of 7%.

# References

# BATC, OMPS SYSTEM DEVELOPMENT for the Ozone Mapping and Profiler Suite (OMPS) Feb. 11, 2009

# BATC Systems Engineering Report, OMPS FM1 Nadir Sensor Pre-Launch Radiometric Calibration Uncertainties, August, 15, 2008

# Fuzhong Weng and OMPS SDR team, Delta Review for SNPP OMPS SDR Earth View Products, August 17, 2015

# BATC, OMPS Nadir Performance & Calibration Test Readiness Review (TRR), December 15, 2006