ATMS Provisional Declaration for TDR, SDR, rSDR

DPA Systems Engineering Evaluation

19 Dec 2012

Product Definitions

- TDR = scene brightness temperature calibrated against warm load and cold space view (a.k.a. antenna temperature)

- SDR = TDR brightness temperature corrected for scan-dependent bias and beam efficiency

- rSDR = SDR brightness temperature CrIS-geolocated and Backus-Gilbert resampled

Requirements Status

- 3 pre-launch performance metrics are described in D46844, Aug 2010, by NG: radiometric calibration, geolocation accuracy, and footprint matching.

- Science team is focusing on 3 requirements of NeDT, calibration accuracy, and geolocation accuracy.

- The SDR has a production requirement phrased in terms of applied corrections

| Req name | Req value | Req reference | Evidence | Product | Comment |
| --- | --- | --- | --- | --- | --- |
| NeDT | Channel-dependent values from 0.5 to 3.6 K | 3 conflicting references: POS, 28300, and STAR | -- Pkg 4 slide 33 says ok -- Pkg 1 slide 8 says ok -- STAR memo says ok -- IDPS single scan average out of spec -- ICVS out of spec | SDR | OPSCON calls for sampling over 100 scans; calc of NeDT used in memo ambiguous |
| Geolocation accuracy | 0.1, 0.2, 0.3 for 3 different beamwidths | OPSCON | Pkg 6 slides 3-16 | TDR |  |
| Calibration accuracy | Channel-dependent values from 0.67 to 2.0 K | 3 conflicting references: POS, 28300, and STAR | Ch 6-13 show Ok in STAR memo; data missing for other channels | TDR | Reflector emissivity issue outstanding |
| Bias & beam efficiency | TDR antenna temperatures are to be corrected for scan bias and beam efficiency | NPP.SDR.3.9 | Not yet provided | SDR | Hooks are in IDPS, but SciTeam must calculate PCT |
| Remapping quality | Not found | Not found | Pkg 5 slides 16-21 | rSDR |  |

Since SDR is built from the TDR, and rSDR from the SDR, issues at a lower level are inherited at the next level.

Several unresolved technical issues have been identified that are relevant to the provisional declaration, and which must be resolved before operational validation:

- TDR issue: striping

- DR 4813; no CCR yet

- Effect of order 0.1 K; within calibration requirements

- Fix will require change to TDR calibration code

- Attributed to warm load noise but mathematical algorithm for correction not determined

- To be corrected for TDR validation

- Science team recommends provisional declaration because effect is within calibration requirements

- TDR issue: reflector emissivity

- No DR filed yet

- Effect of order 1 K for cold scenes (~ 100-200 K), QH at nadir. Marginally at or exceeding calibration requirements

- Manifests as a scan-dependent bias

- STAR proposed compensating in the SDR bias corrections, so no code change is needed for TDR

- Mathematical algorithm  for correction determined, but precise coefficients undetermined

- Science team recommends provisional declaration because correction algorithm is understood, and the effect when viewing earth scenes is within calibration requirements

- SDR issue: beam efficiency and bias correction

- DR 4806; no CCR needed

- EDR PR requirement NPP.SDR.3.9 requires that the SDR contain TDR antenna temperatures corrected for beam efficiency and scan-dependent bias. So far, these corrections have not been applied

- IDPS contains the constructs to allow generation of SDR if the PCT is available. So action is for science team to generate PCT. No IDPS code changes will be needed.

- Science team agrees on beam efficiency calculation approach, but has not agreed on the PCT parameters

- Science team has not agreed on the bias correction

- CrIMSS science team has proposed their own model-dependent bias corrections which will ultimately be coordinated with the bias corrections here

- STAR recommends provisional declaration because “In the provisional presentation package, first three presentations from NWP community, EMC/UKMO/ECMWF, give the TDR data quality preliminary assessment in NWP models”.

- SDR issue: NeDT

- NeDT requirements don’t specify scan averaging

- Single scan averages of the 4 samples exceed requirements. Longer scan averaging can lead to results which meet requirements

- SDR outputs a single scan average which presumably exceed requirements

- ICVS has longer average but still shows requirements exceeded

- NeDT calculation definition which would prove requirement is not in place

- STAR recommends provisional declaration because “There are a few NEdT calculation methods among team and all results are within specification.”

OPSCON Task Status

The ATMS SDR Validation OpsCon document describes 27 cal/val tasks. A final report will be issued for each task in advance of product validation. Recommend providing table of status per task and outlook for validation.