## Slide 1: UW CrIS SDR Status Report
David Tobin, Hank Revercomb, Joe Taylor,
Bob Knuteson, Dan DeSlover, Lori Borg
Suomi NPP SDR Product Review
NCWCP, College Park, MD
23-24 October 2012

## Slide 2
UW CrIS SDR
Cal/Val Tasks
CrIS In-orbit RU Estimation
Internal
consistency checks
on Radiometric Calibration
Radiometric Non-linearity Refinement & Evaluation
Radiometric Noise assessment
Variable
artifact assessment using PCA
Early broadband comparisons with GOES and other GEOs
Clear sky
Obs minus Calc Analysis
Internal consistency checks on spectral calibration, spectral self-apodization correction and resampling
Analysis of non-uniform scene effects on the ILS
SNO comparisons with IASI and AIRS
CrIS/VIIRS Radiance Comparisons
ICT Environmental Model Evaluation and Refinement

## Slide 3: Summary
QC and file latency
  The frequency of repair granules and missing data has decreased substantially over time.
  Users should avoid the associated corrupt spectra by using the QC flags.
Spectral calibration is very good.
  Small changes to the ILS parameters could be considered to further improve the Inter-FOV agreement, most notably for SW FOVs 1 and 6.
CrIS/AIRS comparisons show mean differences of less than 0.1K and stable with time.
  Signal level dependence is very good; some differences at coldest scene temperatures in the SW.
CrIS/VIIRS comparisons show mean differences of less than 0.1K and stable with time.
  Mean agreement improves when the VIIRS OBC is cooled during linearity characterization tests.
  Shows clear dependence on scene BT for M15 @ 10.8um
FOV-2-FOV differences are less than ~30 mK
  Larger differences previously reported for SW regions are an artifact of the analysis technique.
  CrIS/VIIRS comparisons for SW window (and Responsivity variation analysis, not shown) implies good performance of the ICT environmental model.
Nonlinearity looks very good
  a2 changes at the ~5 percent (~30mK) level are being investigated, considered.

## Slide 4: Data QC and File Latency
Corrupt spectra are due to data transmission issues (repair granules, partially full packets). Corrupt spectra include artifacts ranging from a few tenths K to ~100K.
SCRIS file QC Flags based on packet fill percent and Imaginary Radiance components now properly flag/handle the corrupt spectra.
The frequency of repair granules has decreased substantially over time.
However, repair granules are not typically issued within 3 hours, and repair granules do not always appear to incorporated in IDPS/CLASS generated SCRIS files, and so users need to make use of the QC Flags.
Direct Broadcast data and/or data processed using ADL/CSPP after all repair granules have been received do not have these issues.
These comments/results are relative to how UW gets the CrIS data (via IDPS, CLASS, and SD3E). Characteristics of data distributed to DA centers in ~real-time should also be understood.

## Slide 5
Example QC plots for 2012.06.19, IDPS/CLASS products:
Flagged spectra
Real Part
Imaginary Part
Latency plot
Flagged, Case 2
Flagged, Case 3
1. Aggregated RCRIS created at +2 hrs
2. RCRIS repair granules created at +4-5 hrs
3. Aggregated SCRIS files created at +6 hrs

## Slide 6
QC time series
SDRs, data percentage per day
SDRs, good data percentage per day
RDR repair granules per day

## Slide 7: Inter-FOV spectral calibration
The spectral calibration is very good.
Neon lamp views indicate metrology laser variations are less than 1ppm over the last 8 months.
The Inter-FOV calibration is better than 0.2 ppm for LW, 0.3 ppm for LW, 0.7 ppm for SW
Small changes to the ILS parameters could be made to further improve the Inter-FOV agreement, most notably for SW FOVs 1 and 6. (Full spectral resolution data should also be used to assess potential changes).

## Slide 8: Example Center, Side, and Corner FOV ILSs, before Self-Apodization Corrections
Pure sinc
Center FOV5
Edge FOV4
Corner FOV1
 |  | Center FOV | Edge FOV | Corner FOV
centroid (cm-1) | Obs | 942.367 | 942.195 | 942.034
 | Calc | 942.366 | 942.195 | 942.034
FWHM (cm-1) | Obs | 0.747 | 0.757 | 0.767
 | Calc | 0.751 | 0.759 | 0.767
Lfoot | Obs | 0.358 | 0.329 | 0.313
 | Calc | 0.347 | 0.328 | 0.313
Rfoot | Obs | 0.347 | 0.326 | 0.311
 | Calc | 0.345 | 0.329 | 0.313
Calculated
Observed
Notes: The laser wavelength used here and discussed throughout is ½ of the physical wavelength. It is used to represent the distance between samples that is ½ of the laser wavelength because all zeros are used for sampling.

## Slide 9
Metrology laser wavelength deviations, derived from Neon lamp views:
Mean value = 773.13031008 nm
Inter-FOV spectral shifts w/r/t FOV5,
derived from spectral correlation analysis:
Longwave band, 730-740 cm-1
Midwave band, 1350-1380 cm-1
Shortwave band, 2200-2250 cm-1

## Slide 10
Metrology laser wavelength deviations, derived from Neon lamp views:
Mean value = 773.13031008 nm
Inter-FOV spectral shifts w/r/t FOV5,
derived from spectral correlation analysis:
Longwave band, 730-740 cm-1
Midwave band, 1350-1380 cm-1
Shortwave band, 2200-2250 cm-1
Inter-FOV Spectral Cal w/r/t FOV5;
Mean values over last 6 months:

## Slide 11: CrIS/AIRS Radiometric Comparisons
Mean differences are less than 0.1K
Time dependence is very small
Signal level dependence is very good; some differences at cold scene temperatures in the SW.

## Slide 12: CrIS/AIRS dataset, 25 Feb to 7 Oct
455,874 “big circle” samples, 25 Feb to 7 Oct
Scan angles ≤ 30°; Scan angle difference ≤ 3°; Time Diff <= 20 min
AIRS data is L1B v5; CrIS data is ADL (CSPP v1.1) with native Eng. Packets
320K
180K
835 cm-1 BT (K)

## Slide 13: Selected wavenumber regions
To largely avoid the AIRS L1B SRF issues, comparisons shown here are performed for representative ~10 cm-1 regions, selected for sensitivity to CrIS nonlinearity and the CrIS ICT environmental model:
CrIS
AIRS
672-682
830-840
1382-1408
1585-1600
2360-2370
2500-2520

## Slide 14: BT Distributions
CrIS
AIRS

## Slide 15: Daily Mean Differences
V33 Engineering packet upload on April 11 (changes to NL a2 coefficients for LW and MW)
Changes over time are very small.
Largest are for LW 677 cm-1 region, ~30mK, under investigation.
Interesting change in SW 2365 cm-1 behavior in late June.
v33 upload

## Slide 16
-0.061 ± 0.001 K
0.020 ± 0.004 K
-0.077 ± 0.002 K
0.067 ± 0.002 K
0.038 ± 0.006 K
-0.058 ± 0.003 K
BT Difference Distributions, 13 Apr – 7 Oct

## Slide 17: CrIS/VIIRS Radiometric Evaluations (in collaboration with Chris Moeller)
The comparisons are beneficial for both VIIRS and CrIS
Mean differences are <0.1K
Comparisons shows clear dependence on scene BT for M15, with VIIRS cooler than CrIS by 0.4K at 205K.
Mean agreement improves when the VIIRS OBC is cooled during quarterly linearity characterization tests.
Small changes in mean differences (~12mK) over last 6 months.

## Slide 18
Monochromatic spectrum, CrIS spectrum, and VIIRS SRFs
BT (K)
wavenumber, 1/cm
BT (K)
wavenumber, 1/cm
M13
4um
M15
10.8um
M16A
12um

## Slide 19
CrIS processing is CSPP with v33 Eng. Packet; VIIRS is IDPS product
Each day includes ~500,000 colocations which pass a spatial uniformity test
Daily mean differences are < 0.1K since VIIRS OBC LUT change in early March
Major discontinuities are due to known events (e.g. VIIRS OBS LUT change in early March, shutdown/restart on March 24/25, VIIRS OBC temperature ramp in late May and mid Sept)
Slow trends in all three bands, ~12 mK, since May

## Slide 20
mean scene BTs
M13 differences show little dependence on scene BT, except for coldest scenes
M15 and M16 show clear scene BT dependence of differing magnitude

## Slide 21
As the VIIRS OBC temperature approaches the instrument temperature, the VIIRS calibration becomes less sensitive to knowledge of the OBC emissivity and to knowledge of the instrument temperatures, and also changes the nonlinearity “set point”; Changes to the CrIS/VIIRS comparisons during this test implies that small improvements to the VIIRS calibration parameters are possible.
For warm scenes, cooling the OBC temperature causes (1) M13 biases to converge with those of M15 and M16, as well as (2) better overall agreement with CrIS.
Figure c/o
C. Moeller

## Slide 22: FOV-2-FOV Radiometric Differences
FOV-2-FOV differences are less than ~30 mK
Larger differences previously reported for SW band are an artifact of the analysis technique.
CrIS/VIIRS comparisons for SW window implies good performance of the ICT environmental model.

## Slide 23
IDPS SCRIS files
LW, 672-682 cm-1, wrt FOV5
LW, 830-840 cm-1, wrt FOV5
MW, 1382-1408 cm-1, wrt FOV9
MW, 1585-1600 cm-1, wrt FOV9
SW, 2360-2390 cm-1, wrt FOV5
SW, 2500-2520 cm-1, wrt FOV5
“CrIS-only” FOV-2-FOV differences;
Daily mean differences

## Slide 24
“CrIS-only” FOV-2-FOV differences;
Mean Differences over last 6 months
LW and MW values are less than ~30mK.
SW values are larger, particularly for FOVs 3,6, and 9, for both opaque and window regions.

## Slide 25: CrIS/VIIRS M13 (4um) comparisons
CrIS/VIIRS comparisons, broken out by CRIS FOV, offer an independent way to assess CrIS FOV-2-FOV differences; Less impact of view angle differences present in CrIS-only approach.
SW window region is most sensitive to the CrIS ICT environmental model, which includes the modeled term for the SSM Baffle which varies with orbit phase.
Variations shown here are well behaved as a function of orbit phase. Some low level behavior (~0.1K) seen for FOVs 7 and 8.

## Slide 26
The CrIS/VIIRS comparisons for SW M13, opposed to the CrIS-only results, do not show the larger differences for FOVs 3,6,9.
The CrIS/VIIRS comparisons for LW M16 show very good agreement with the CrIS-only results.
FOV-2-FOV differences for all bands/regions are less than ~30mK
Currently investigating impacts of CrIS-only approach limitations on other wavelength regions.
“CrIS-only” and CrIS/VIIRS derived FOV-2-FOV differences

## Slide 27: Radiometric Nonlinearity
NL corrections look very good, based on FOV-2-FOV and other analyses.
a2 changes below the ~10 percent level are being investigated. (Pre-launch characterization had 10% (LW) and 15% (MW) 1-sigma uncertainty).

## Slide 28
NLC: C’ = C / (1 - a2 VDC )
Example Impact of Nonlinearity Correction on calibrated Earth view radiances
Notes: Warm scene. ~clear sky
Positive and negative signatures in LW.

## Slide 29
v32 (TVAC, yellow) and v33 (In-orbit, orange) a2 values
Notes:   NL Summary:
  Overall, the FOV-2-FOV analysis is working well.
  MW7 value changed by ~100% (not unexpected)
  MW6 and 9 remained linear
  Overall, FOV-2-FOV values are larger (~10%) than v32 values.
  But does not necessarily suggest actual changes in all FOVs due to differences in TVAC vs. In-orbit analysis uncertainties
  SW biases are not zero, and are FOV dependent.
  The next iteration of Nonlinearity analysis will likely require modifications to the analysis technique, addressing the SW biases and other possible sources of radiometric bias.

## Slide 30
Nonlinearity Monitoring
Create daily estimates of fractional change in detector a2 values.
Example below for FOV-2-FOV differences shown previously.
No conclusive evidence for a2 change exceeding the 10% level.

## Slide 31: October 6 Suomi-NPP underflight
A “piggy-back” flight, funded by JPSS Program Science, as part of the recent NASA HS3 campaign.
A dedicated Suomi-NPP campaign will be held in May 2013 in support of the "Validated" SDR review.

## Slide 33
Preliminary
S-HIS/CrIS Radiance comparison:

## Slide 34: Summary
QC and file latency
  The frequency of repair granules and missing data has decreased substantially over time.
  Users should avoid the associated corrupt spectra by using the QC flags.
Spectral calibration is very good.
  Small changes to the ILS parameters could be considered to further improve the Inter-FOV agreement, most notably for SW FOVs 1 and 6.
CrIS/AIRS comparisons show mean differences of less than 0.1K and stable with time.
  Signal level dependence is very good; some differences at coldest scene temperatures in the SW.
CrIS/VIIRS comparisons show mean differences of less than 0.1K and stable with time.
  Mean agreement improves when the VIIRS OBC is cooled during linearity characterization tests.
  Shows clear dependence on scene BT for M15 @ 10.8um
FOV-2-FOV differences are less than ~30 mK
  Larger differences previously reported for SW regions are an artifact of the analysis technique.
  CrIS/VIIRS comparisons for SW window (and Responsivity variation analysis, not shown) implies good performance of the ICT environmental model.
Nonlinearity looks very good
  a2 changes at the ~5 percent (~30mK) level are being investigated, considered.

## Slide 35: Future Efforts
  “Spectral ringing”
    Currently investigating with PCA, Earth view DM data, and the possible role of the NF
  Inter-FOV spectral calibration
    Considering small changes to the ILS parameters
  Nonlinearity
    Evaluating small changes to a2 at the ~5% level
    Incorporate other evaluations (e.g. obs-calc, CrIS/AIRS/IASI intercal, revisit DM analysis, etc.)
  Full Spectral resolution
    Re-visit ILS and Nonlinearity assessment when in full resolution mode
  Non-uniform scene effects on the ILS
  Aircraft Campaign participation with the Scanning-HIS
  Overall RU estimation
    Include refined estimates of contribution uncertainties, and new terms as needed, to produce an updated model of CrIS Radiometric Uncertainty