## Slide 1: Evaluation of OMPS sensor and EV SDR performance
Conclusions:
Sensors are performing well
Calibrations are very good; no sign of degradation
Comparisons between IDPS & PEATE SDRs exhibit troubling differences
Differences have little effect on ozone, but cause us to question IDPS SDR
We recommend AGAINST Provisional status at this time
NASA’s OMPS Ozone Science Team
NASA Ozone Product Evaluation and Algorithm Test Element (PEATE)
JPSS OMPS Science Operations Center (SOC)

## Slide 2
Total Column
Nadir Profiler
A “hot pixel” is any pixel with a statistically significant increase in dark current since launch
Increasing dark currents, as expected

## Slide 3
Knowledge spec. is ±0.2% over the full dynamic range
Sensor linearity is stable
NP Linearity Measurement regression residuals
Orbits 230 - 1166
Qideal / 1000
Full well

## Slide 4
Neither Nadir sensor has exhibited signs of degradation since first light measurements in January, 2012 except possibly at the very shortest wavelengths. Results also indicate that seasonally dependent sensor response corrections are very accurate.
Solar measurements show no evidence of sensor change

## Slide 5
The wavelength registration is derived by comparison with a synthetic solar spectrum based on ATLAS SUSIM and Kitt Peak (also used for Aura/OMI).
OMPS TC Measured
Reference Spectrum
OMPS NP Measured
Reference Spectrum
Wavelength (nm)
Wavelength (nm)
TC & NP wavelengths have been adjusted using measured solar spectra
Shift results suggest an accuracy of ~1/100 pixel.

## Slide 6
Chart: Wavelength [nm]
OMPS Day 1 Solar Flux compared to synthetic spectrum after wavelength shift
OMI comparison to same synthetic spectrum
Solar spectra compare well with previous extraterrestrial measurements

## Slide 7
OMPS Swath Coverage on VIIRS Image
(yellow=TC, red=NP)
Nadir Mapper (TC) TOA Refl.
310.70 nm Band (33)
Nadir Profiler (NP) TOA Refl.
310.73 nm Band (35)
Colocation of TC & NP is reasonably good
OMPS EV_COLOC Orbit 1616 over Western Australia on 2/19/2012

## Slide 8
Co-locate MLS temperature, ozone profiles to OMPS TC measurements
  Reflectivity < 0.10
  -10 < latitude < 10 degrees
  June 2012
Calculate TOA reflectances (radiance / solar flux) from TC viewing conditions, MLS profiles using radiative transfer code (TOMRAD)
Compare measured OMPS TC reflectance with calculated reflectance
  Agreement seen to within 1% for wavelengths > 312 nm
Stray light seen for wavelengths < 312 nm
  Consistent with pre-launch sensor characterization
Stray light
Structure seen in spectrum
repeats in all swath positions,
taken out through dividing by
position 18 (Raman scattering)
Position 19 (nadir looking)
Position 19 / Position 18
Percent reflectance difference (meas – calc)
Wavelength (nm)
OMPS TC comparison with modeled TOA reflectances is quite good

## Slide 9
Wavelength (nm)
Position 36 / Position 18
Position 33 / Position 18
Difference indicates
calibration issue
Position 4 / Position 18
Position 1 / Position 18
Percent NR difference (meas – calc)
Difference indicates inconsistency
with MLS ozone profile, will lead to
different total column ozone amounts
Difference indicates
calibration issue
Difference indicates inconsistency
with MLS ozone profile, will lead to
different total column ozone amounts
OMPS TC cross-track calibration is typical
Problems at the far off-nadir positions lead to swath dependent ozone effects
Not unusual

## Slide 10
331 nm Reflectivity
Radiance / cos(SolZA)
SBUV2 technique for stray light evaluation applied to NP
+
+
+
+
+
+
Model estimate
Measurement
Orb 2840
Pre-launch stray light model applied to real data
NP stray light appears “as expected”
Errors in 1 mb O3 could be 5% or more

## Slide 11
Assume NP radiances contain no stray light in the overlap region (300 – 310 nm)
Use TC / NP comparison to tune ghost
Best result yields ±5% residual errors at 300 nm (except at high SolZA)
NP / TC Radiance Ratio
TC stray light can be corrected
TC Stray Light is partly caused by a ghost from 417 nm – not well characterized pre-launch
No Correction
Basic Correction
Basic Correction + various ghost models
Orbit 2840

## Slide 12
OMPS Science Operations Center
O3 PEATE
NGAS
C3S
IDPS
SD3E
NASA Cal. SDR
OMPS RDRs
Analysis Results
LUTs
Fast Track & Diag. Flight Tables
OMPS Operations
The OMPS SOC utilizes the NASA Cal. SDR product to maintain Flight and Ground calibrations
Flight calibration operations were modified after launch to optimize measurements

## Slide 13
Table Description | Table Type | Delivery Status
TC & NP Day 1 Solar | LUT | Once (repeat is TBD)
TC & NP Wavelength | GND-PI | Monthly (ceased with Mx6.3)
TC & NP CF Earth | GND-PI | Monthly (ceased with Mx6.3)
TC & NP Dark Tables | GND-PI | In test
Diagnostic Flight Sample Tables | SCT | Frequently
Earth-view Flight Sample Tables | SCT | Once
Earth-view Ground Sample Tables | GND-PI | Once
Calibration Flight Sample Tables | SCT | Once
TC & NP Radiometric Coefficients | LUT | TBD
TC Stray Light Coefficients | LUT | In test
NP Stray Light Coefficients | LUT | Not planned
TC & NP Linearity (Flight & Ground) | SCT/GND-PI | Not planned
TC & NP Flat Field | SCT | Not planned
OMPS SOC delivers Cal. tables

## Slide 14
TC Solar (1-4) | 12 images
NP Solar | 9 images
TC Solar (5-7) | 9 images
NP FF Lamp | 1 image
NP Linearity | 83 images
NP FF Dark | 1 image
NP Dark coadd | 1 image
TC FF Lamp | 1 image
TC Linearity | 83 images
TC FF Dark | 1 image
TC Dark coadd | 1 image
EVNCAL – 1 orbit
Every week
TC Solar (1,4,7) | 57 images
TC / NP Stor. Darks | 37 images
TC Solar 2 | 16 images
NP Solar | 16 images
TC Solar 6 | 16 images
TC / NP Open Darks | 37 images
TC Solar (1,3,5,7) | 54 images
TC / NP Closed Darks | 37 images
3orb Solar
Every 2nd week
Orbit 1
Orbit 2
Orbit 3
NP Lamp Warmup | 50 images
NP Linearity | 83 images
NP FF Lamp | 1 image
TC / NP Closed Darks | 5 images
TC Lamp Warmup | 50 images
TC Linearity | 83 images
TC FF Lamp | 1 image
EVLED_Closed – 1 orbit
Every 4th week
TC / NP Closed Darks | 21 images
TC / NP Stor. Darks | 9 images
EV_4WK_DC – 1 orbit
Every 4th week
Flight calibrations modified after launch
Original
Modified

## Slide 15
Original dark sequence was not optimized for transient removal
Updated dark calibration sequence
Before transient removal
After transient removal
Dark Correction for 7.5 sec TC Image
(row by row mean)
PEATE / SOC Darks
Original Cal. Darks
We do not currently understand this small dark current difference

## Slide 16
Solar noise driven by diffuser features
NP Working Diffuser ratio: Aug. 31 / Mar. 21
  β=18.95° / β=18.95°
NP Working Diffuser ratio: Jun. 27 / Mar. 21
  β=14.57° / β=18.95°
Solar β Angle (degrees)
σ of ratio to Mar. 21 (%)
NP
TC

## Slide 17
Special solar data resampled to study diffuser feature reduction
3-image running aggregation
12-image running aggregation
Original solar sequence
Original solar sequence
New solar sequence: ~ 80 sec
Solar sequence lengthened to reduce noise from diffuser features

## Slide 18
NASA evaluation of IDPS OMPS SDR
What does the SDR do?
Radiance = [ Raw Counts – Bias – Smear – Dark – StrayLight ] x Cal. Consts.
OK
Problems
Not implemented
Problems
OK?
Smear, Dark, and StrayLight are small corrections, so radiance errors are small
Some problems have clear resolutions
For others the cause is still under investigation
  EV SDR not understood well enough for Provisional
But …. we believe this evaluation is about establishing confidence in SDR

## Slide 19
Transients in smear region not corrected
Smear is the only correction derived directly from Earth scene data
Smear Signal
Derived Radiance (right side)
Derived Radiance (left side)
These data are doors-closed, so radiances should be zero
Will result in significant radiance errors over entire swath (primarily in SAA)

## Slide 20
Do not understand how
these values come about
given the existing co-add and binning
characteristics
Swath corresponding to
12 Oct 2012
5 hr, 52 min, 2.82 sec
PEATE SDR
Bias Floor
smear
IDPS SDR
smear
Wavelength Index
Wavelength Index
Smear values (reported as used in SDR) don’t appear to be correct

## Slide 21
IDPS SDR uses radiance conversion coefficients that already assume
that a stray light correction has already been applied to the radiances
No such correction has yet been applied
FOV 16
IDPS
PEATE
IDPS / PEATE
RadCoef SLcor / RadCoef NOcor
Calibration constants not appropriate for current stray light correction

## Slide 22
Code doesn’t properly deal with SAA corrections
IDPS
PEATE
SDR code currently uses a “special” Dark subtraction within the SAA
Smear values used by SDR are all negative within the SAA
OMPS doesn’t produce negative numbers

## Slide 23
SOC provided Darks GND-PI table
G-ADA SDR run DarkCurrentEarth
Special dark table generated to study interpretation by EV SDR code
Test confirms that the EV SDR assigns darks to pixels the same way the SOC does when creating the Dark GND-PI table

## Slide 24
Expected Darks in radiance output
Special test designed so that only dark counts show up in radiance output
We recently discovered that the test was ill-posed for testing darks. But these results are consistent with a smear problem.
Observed Darks in radiance output

## Slide 25
Shameless website plug
http://ozoneaq.gsfc.nasa.gov/omps

## Slide 26
Backups

## Slide 27
MLS Comparisons

## Slide 28
The accuracy of the Nadir Mapper and Profiler calibration was assessed by comparing measurements with radiative transfer model calculations for top-of-the-atmosphere reflectance. These calculations are based on coincident ozone profile retrievals from the NOAA-18 SBUV2 instrument.
Mapper Cross-track Calibration
Reference wavelengths
Nadir Calibration Error
Nadir Mapper
Nadir Profiler
Radiative transfer calculation error ?
Calibration adjustments applied to PEATE products

## Slide 29: Radiance difference divided by Radiance coefficient
Expected Darks4
Observed Darks4
The left slide is same as that in previous slide but plotted with the same axis scales on the right