## Slide 1
The ACOS/GOSAT Experience
David Crisp JPL/CALTECH
Hyper Spectral Workshop, Miami
Question 8: Japan’s Greenhouse Gases Observing Satellite, GOSAT, and its IR hyperspectral sensor, TANSO-FTS, measuring CO2 and CH4:
  How well has it performed?
  What are the experiences with its operations?

## Slide 2: OCO and GOSAT Collaboration
The OCO and GOSAT teams formed a close partnership during the implementation phases of these missions to:
  Cross calibrate the OCO instrument and TANSO-FTS
  Cross validate OCO and GOSAT XCO2 retrievals against a common standard
The primary objectives of this partnership were to:
  Accelerate understanding of this new data source
  Facilitate combining results from GOSAT and OCO
3-day ground track repeat cycle resolves weather
Continuous high resolution measurements along track
GOSAT
OCO

## Slide 3: The Launch of GOSAT & Loss of OCO
24 Feb 2009
GOSAT launched successfully on 23 January 2009
OCO was lost a month later when its launch system failed
Notes: OCO is designed to provide the data needed to find the missing pieces of the carbon puzzle (sinks)

## Slide 4: Working with the GOSAT Team
Immediately after the loss of OCO, the GOSAT Project manager invited the OCO Team to participate in GOSAT data analysis
NASA reformulated the OCO team as the “Atmospheric CO2 Observations from Space” (ACOS) team
This collaboration benefits the GOSAT team by:
  Combining the ground based calibration and validation resources of both teams to maximize the accuracy of the GOSAT data
  Combining the scientific expertise from both teams to accelerate our understanding of this new, space-based data source
This collaboration benefits the NASA OCO by
  Providing direct experience with the analysis of space based CO2 measurements
  Accelerating the delivery of precise CO2 measurements from future NASA carbon dioxide monitoring missions

## Slide 5: Scope of the ACOS/GOSAT Collaboration
The ACOS team is collaborating closely with the GOSAT teams at JAXA and NIES to:
  Conduct vicarious calibration campaigns in Railroad Valley, Nevada, U.S.A. and analyze results of those campaigns
  Retrieve XCO2 from GOSAT spectra
      Model development, implementation, & testing
      Data production and delivery
  Validate GOSAT retrievals by comparing
      GOSAT retrievals with TCCON measurements
      Other validation standards (surface pressure, aircraft and ground-based CO2 measurements
Picarro CO2

## Slide 6: Retrieving XCO2 from GOSAT Data
Forward Model
Spectra + Jacobians
Calibrated GOSAT Spectra
(L1B Data)
State Vector First Guess
Update State Vector
Inverse Model
(Optimal Estimation)
Apriori + Covariance
Calculate XCO2
Diagnostics
converged
not converged
State Vector
CO2 profile (full)
H2O profile (scale factor)
Temperature profile (offset)
Aerosol Profiles
Surface Pressure
Albedo (Mean, Slope)
Wavelength Shift (+ stretch)
The OCO Retrieval Algorithm was modified to retrieve XCO2 from GOSAT measurements
“Full-physics” forward model
Inverse model based on optimal estimation

## Slide 7: ACOS/GOSAT Data Products: XCO2 Retrievals
XCO2 Retrievals for July 2010

## Slide 8: ACOS/GOSAT Data Products: Surface Pressure Retrievals
Surface Pressure Retrievals for July 2010

## Slide 9: ACOS/GOSAT Data Products: Aerosol Optical Depth Retrievals
AOD Retrievals for July 2010

## Slide 10: Experiences with its operations: The ACOS Cloud Screen
A Spectroscopic cloud screening algorithm based on the O2 A-band is currently being used for GOSAT retrievals
  Fits a clear sky atmosphere to every sounding in the O2 A band.
  High values of 2 and large differences between the retrieved surface pressure and the ECMWF prior indicate the presence of clouds
  Over non-glint ocean, a simple albedo test is also used.
Example A-Band fit
Poor fit (2 = 9.6) indicates presence of cloud
Small residuals and good agreement between retrieved and ECMWF surface pressure indicates cloud free

## Slide 11: Experiences with its operations: Post Screening Improves Accuracy
Errors can be further reduced by post-screening retrievals, based on a series of criteria, including:
Measurement SNR
Convergence
Goodness of spectral fit
Surface pressure error
Evidence for clouds or optically thick aerosols
A postiori retrieval error
Evidence of known biases
The cloud screen is responsible for the largest data reductions.
Improved cloud screening algorithms are a major focus of our development effort

## Slide 12: Experiences with its operations: Validation of GOSAT Products
447-m WLEF Tower
Park Falls, WI
FTS
GOSAT XCO2 retrievals are being compared with those from the ground based Total Carbon Column Observing Network to very their accuracy
Current and Planned TCCON Stations

## Slide 13: How Well has it Performed? Comparisons of GOSAT and TCCON
ACOS GOSAT retrievals show
  A consistent global bias of ~2% (7 ppm) in XCO2 when compared with TCCON and aircraft measurements.
  XCO2 variations that are a factor of 2 to 3 larger than that measured by TCCON.
Wunch et al.

## Slide 14: Sources of Bias in the XCO2 Maps
About half of the global 2% bias is caused by a ~10 hPa (1%) high surface pressure bias associated with:
  Radiometric and spectroscopic calibration errors
    Non linearity and ILS corrections currently being implemented
  Uncertainties in the O2 A-band absorption cross sections
    Oversimplified treatment of line mixing and line shape
Improvements in the CO2 spectroscopy and aerosol retrieval approach are also being implemented to address
remaining bias
Ps – Ps (a priori)
Typical O2 A-band retrieval residuals.
GOSAT

## Slide 15: Validation of XCO2 at TCCON Sites
Once the known biases are removed, retrievals of XCO2 compare well against “ground truth” from the TCCON (Total Carbon Column Observing Network).
Both the hemispheric gradients and the seasonal cycles are captured in the bias-corrected GOSAT XCO2 retrievals.
Park Falls, WI
Lamont, OK
Lauder, NZ

## Slide 16: Experiences with its operations: A Year of ACOS/GOSAT XCO2

## Slide 17: ACOS GOSAT Data Release
The ACOS L2 Standard Products are now being distributed on the GSFC Mirador site.
http://mirador.gsfc.nasa.gov/
You can find the data, along with a README and a Data Quality Statement at the site:
http://disc.sci.gsfc.nasa.gov/acdisc/documentation/ACOS.shtml

## Slide 18: Conclusions and Future Plans
The ACOS/GOSAT collaboration is returning benefits to both teams
  The OCO-2/ACOS retrieval algorithm is:
    currently in place and is generating a production product for GOSAT
    still evolving, to address know errors and biases
  The vicarious calibration experiments have helped to identify and correct for changes in the pre-launch GOSAT calibration parameters.
  Comparisons with TCCON measurements have revealed a global, -2% bias in the preliminary ACOS XCO2 retrievals
    About 2/3 of this bias can be attributed to a +10 hPa bias in the retrieved surface pressure
Lessons learned from this experience are expected to substantially accelerate the delivery of high quality products from the OCO-2 mission, which is currently scheduled for launch in February 2013